## Slide 1
38655 BMED-2300-02
Lecture 13: CT Reconstruction
Ge Wang, PhD
Biomedical Imaging Center
CBIS/BME, RPI
wangg6@rpi.edu
March 6, 2018

## Slide 2: Are Your Scores Good?

## Slide 3
Tue | Topic | Fri | Topic
1/16 | Introduction | 1/19 | MatLab I (Basics)
1/23 | System | 1/26 | Convolution
1/30 | Fourier Series | 2/02 | Fourier Transform
2/06 | Signal Processing | 2/09 | Discrete FT & FFT
2/13 | MatLab II (Homework) | 2/16 | Network
2/20 | No Class | 2/23 | Exam I
2/27 | Quality & Performance | 3/02 | X-ray & Radiography
3/06 | CT Reconstruction | 3/09 | CT Scanner
3/20 | MatLab III (CT) | 3/23 | Nuclear Physics
3/27 | PET & SPECT | 3/30 | MRI I
4/03 | Exam II | 4/06 | MRI II
4/10 | MRI III | 4/13 | Ultrasound I
4/17 | Ultrasound II | 4/20 | Optical Imaging
4/24 | Machine Learning | 4/27 | Exam III
BB Schedule for S18
Office Hour: Ge Tue & Fri 3-4 @ CBIS 3209 | wangg6@rpi.edu
Kathleen Mon 4-5 & Thurs 4-5 @ JEC 7045 | chens18@rpi.edu

## Slide 4: Some Tips
Basic Books:
Ge’s Book & Green Book
Understand Logic Flow/Key Ideas from Lectures/PPTs
Gain Skills from Homework & Textbook Questions
Have Habit of Doing Preview-Listening-Review, & Repeated Reviews

## Slide 5: CT Reconstruction
Duality of Information
Algebraic Approach
Solution Uniqueness
Data Independence/Sufficiency
Analytic Approach
Fourier Slice Theorem
Filtered Backprojection
Examples
Numerical Examples
Clinical Images

## Slide 6: Duality of Information
Particle
Wave
Image = Collection of Pixels/Voxels
Image = Superposition of Waves/Wavelets

## Slide 7: Pixel: Picture Element

## Slide 8: Fourier Analysis

## Slide 9: CT Reconstruction
Duality of Information
Algebraic Approach
Solution Uniqueness
Data Independence/Sufficiency
Analytic Approach
Fourier Slice Theorem
Filtered Backprojection
Examples
Numerical Examples
Clinical Images

## Slide 10: Linear Equation
x
No
Ni
k
Ray-sum
Line integral
X-ray measurement gives data as ray sums or line integrals

## Slide 11: Data Independence
By “onion-peeling”, each x-ray path carries unique information, and all “pixels” can be resolved.

## Slide 12: Data Sufficiency
There exists at least a source on any line
intersecting a cross-section
These represent the maximum amount of data you can measure, which happens to be sufficient

## Slide 13: Projection & Sinogram
Sinogram
t

Sinogram (Radon Transform): All projections
p(t)
f(x,y)
t

y
x
X-rays
Projection: All ray-sums in a direction
p
p(t)

## Slide 14: Computed Tomography
p(t)
f(x,y)
Computed tomography (CT): Image reconstruction from projections
p(t)
f(x,y)
t

y
x
X-rays

## Slide 15: Sum Game
=4
2=3
3=2
4=1
A practical problem has many more unknowns but the idea is the same!

## Slide 16: Trial & Error
3
2
3
2
4
3
2
1
0
0
0
0
4
3
2
1
6
4
-2
2
Update a guess based on data differences
Guess 1
Guess 0
Guess 2
Error
Error

## Slide 17: Algebraic Approach
Convert X-ray data into line integrals to form a system of linear equations
Solve the system of linear equations to reconstruct an underlying image
Regularize the image reconstruction with various prior knowledge
Iteratively refine an intermediate image until the outcome is satisfactory

## Slide 18: CT Reconstruction
Duality of Information
Algebraic Approach
Solution Uniqueness
Data Independence/Sufficiency
Analytic Approach
Fourier Slice Theorem
Filtered Backprojection
Examples
Numerical Examples
Clinical Images

## Slide 19: Probing Waves
No any info in projection data along non-orthogonal directions
Info in projection data along the orthogonal direction

## Slide 20: Fourier Slice Theorem
2D Fourier Transform
1D Fourier Transform
X-rays
Imagine that f(x,y) is a superposition of various 2D waves propagating along different orientations at all frequencies… …

## Slide 21: Kak’s Book Chapter 3

## Slide 22: Simplest Example

## Slide 23: Along Vertical Direction

## Slide 24: Along Any Direction

## Slide 25: Fourier Slice Theorem

## Slide 26: Fourier Slice Theorem

## Slide 27: From Projections to Image
y
x
v
u
F-1[F(u,v)]
f(x,y)
P(t)
F(u,v)
X-ray data are essentially samples in the Fourier space, and image reconstruction can be done via inverse Fourier transform!

## Slide 28: Filtered Backprojection
The idea is to express f(x,y) in terms of the Fourier transform in polar coordinates

## Slide 29: Filtered Backprojection
Half-scan over 0-180 Degrees Gives Sufficient Information for Image Reconstruction

## Slide 30: Filtered Backprojection

## Slide 31: Ramp Filter

## Slide 32: Bandlimited Expression

## Slide 33: Digital Filtration

## Slide 34: 2D Fourier Transform
3D Fourier Transform
Extended Fourier Slice Theorem
Fourier slice theorem can be generalized into other forms… …

## Slide 35: Analytic Approach
Convert X-ray data into the Fourier/Radon space
Invert the Fourier/Radon transform according to a closed-form formula
In the reconstruction process, some filtering/processing steps may be used
Noise/incompleteness/inconsistency may induce stronger artifacts than iterative reconstruction

## Slide 36: CT Reconstruction
Duality of Information
Algebraic Approach
Solution Uniqueness
Data Independence/Sufficiency
Analytic Approach
Fourier Slice Theorem
Filtered Backprojection
Examples
Numerical Examples
Clinical Images

## Slide 37: Image to Projection
Sinogram
Projection
Image

## Slide 38: Backprojection
Projection

## Slide 39: Backprojection to Image
Sinogram
Backprojected Image

## Slide 40: Filtered Backprojection
f(x,y)
f(x,y)
P(t)
P’(t)
1) Convolve projections with a filter
2) Backproject filtered projections

## Slide 41: Filtered Backprojection

## Slide 42: Projection to Filtered Projection
Filtered Sinogram
Sinogram

## Slide 43: Filtered Backprojection to Image
Filtered Sinogram
Reconstructed Image

## Slide 44: Radon Transform

## Slide 45: Radon Parameters

## Slide 46: Parallel Beam Geometry

## Slide 47: Projection Algorithm

## Slide 48: Sample Radon Code
I = zeros(100,100);
I(25:75,25:75) = 1;
imshow(I)
theta = 0:180;
[R,xp] = radon(I,theta);
imagesc(theta,xp,R);
title('R_{\theta} (X\prime)');
xlabel('\theta (degrees)');
ylabel('X\prime');
set(gca,'XTick',0:20:180);
colormap(hot);
colorbar

## Slide 49: Sinogram Display

## Slide 50: Sample iRadon Code
P = phantom(128);
R = radon(P,0:179);
I1 = iradon(R,0:179);
I2 = iradon(R,0:179,'linear','none');
subplot(1,3,1), imshow(P), title('Original')
subplot(1,3,2), imshow(I1), title('Filtered backprojection')
subplot(1,3,3), imshow(I2,[]), title('Unfiltered backprojection')

## Slide 51: Reconstructed Results

## Slide 52: Clinical Examples

## Slide 53
BB13 Homework
Read the web page carefully and run the codes on your PC:
http://www.mathworks.com/help/images/radon-transform.html?refresh=true
Then write a similar page with an ellipse as your object (instead of the square).
Use the iradon function to reconstruct your ellipse.
Due Date: Same (Week Later)